Take with you: Specimen (In transfer container with LID due to no ventilation for formalin)

    - Gloves

    - Paperwork or patient sticker

      - (need name, MRN, case number, and laterality)

    - T-pins

    - Room key from wall in accessioning area

      - (door stays locked-- don’t forget the key in the room!)

Transporting specimen:  Needs to be in closed and cleaned container as you are walking in “clean” areas and public hallways.

  - No gloves on door handles/ remove PPE prior to leaving the lab

  - Keep computer keyboard and mouse clean (only handle with clean gloves)

DON’T FORGET TO SAVE THE IMAGE (SEE BELOW)!!

Wipe the machine, stage, work area and keyboard down with cavi wipe when you are finished

How to use the Faxitron:

Power-up: The unit should be up and running, but if not, there is a power switch on the back near the power cord. Also, make sure the computer is on (power on CPU)

Start Software:  Double click on Vision Icon on desktop

If calibration needs to be done- allow 5 minutes and when the status bar stops advancing, click pause, and then click play/resume and it should finish.

Getting Started: Start procedure  Patient Data  Manual Entry

Enter required (*) information along with case number followed by part type (ie S18-88453b)

X-Ray Imaging: Place specimen in center of tray noting orientation. Standard orientation is superior towards back of unit and anterior placed on the tray. Click start x-ray or push x-ray button on front of unit.

Image Manipulation: Hold left mouse button and drag up/down for brightness and left/right for contrast

Click full resolution and drag hand in window for magnified view

Invert image to change black to white

Saving Image: If toolbar is not visible click on ᴠ symbol on the top right of the screen

Shut-Down: Clean unit and counter tops with cavi wipes and leave unit and computer running for next person.